## Slide 1
Advanced Instrumentation for Ion Nano-Analytics
Materials Research and Technology Department
Luxembourg Institute of Science and Technology
SIMS imaging performed on Focused Ion Beam - based platforms
Jean-Nicolas Audinot,
A. D. Ost, C. Stoffels, P. Philipp, O. De Castro, A. Biesemeier, Q. H. Hoang and T. Wirtz
Notes: First I’ll like thank the organizers to give me the opportunity to present our research ‘ work done , and on going, at LIST

## Slide 2: Focused Ion Beam platform
Applications Areas
Materials Science
Nanofabrication
Industrial Research
Life Sciences
Circuit Editing
Oil and Gas
Lithography
Nanophotonics
Polymer Imaging
Polymer Imaging
Lipid nanodomains
Nematode
Scientific Reports volume 3, Article number: 3514 (2013)
Small 2015, 11, No. 43, 5781–5789
FIB
FIB

## Slide 3: Ionic Source : State of art
 | Electron impact ion source | Duoplasmatron ion source | Surface ionisation ion source | Hyperion plasma source | Liquid metal ion source | Gas field ionisation ion source | Cold atom ion source
Species | Ar, Xe, O2, SF6, C60, … | Ar, Xe, O2, … | Cs | H, He, O2, Ar, Ne, Kr, Xe | Ga, In, Aun, Bin, … | He, Ne, N2 | Li, Na, Rb, Cs, Ga, …
Virtual source size | 1 mm | 200 µm | 50 µm | 15 µm | 3 nm | ~3 Å | 10 µm
Ion energy for singly charged ions (keV) | 0.1–20 | 0.1–20 | 0.3–20 | 5–30 | 2–30 | 5–45 | 2 – 30
Beam current | 1–1000 µA | 1–1000 mA | 10 µA–100 mA | 1 pA - 1 µA | 1–1000 µA | 1 fA–100 pA | 1 pA – 10 nA
Brightness (A cm-2 sr-1) | 1 | 102 | 102 | 103 | 107 | 109 | >107
Energy dispersion (eV) | 1–10 | 5–20 | 0.2–0.5 | 5–6 | 5–50 | <1 | 0.01 - 10
New FIB platforms
Audinot et al. Reports on Progress in Physics (2021)

## Slide 4: FIB platform
Ga Liquid Metal Ion Sources
Gas Field Ion Source
Low-Temperature Ion Source
0,5 nm / 2 nm
2 nm
2 nm
He+ / Ne+
Ga+
Cs+
Analytical capabilities limited
no X rays emission (EDX, WDS,… )
but ion emission ……….

## Slide 5: Surface analysis Techniques
TEM
EELS/EDS
AES
LA-ICP-MS
RAMAN
SEM EDS
XPS
XRD
XRF
FTIR
SIMS
Tof-SIMS
µ-XRF
µ-PIXE
100 %
10 %
1 %
0.1 %
100 ppm
10 ppm
1 ppm
100 ppb
10 ppb
1 ppb
100 ppt
0.1 nm
1 nm
10 nm
100 nm
1 µm
10 µm
100 µm
1 mm
1 cm
Nano
Micro
Bulk
ICP-MS
Lateral resolution
Detection limit
Physical limit for 0,3 nm sampling depth
NanoSIMS
TEM
AFM
SEM/HIM
GDMS
TGA/DTA/DSC
ICP Techniques
GC-MS
Optical
Spatial resolution
Incident Particle
X Ray
X Ray
h
h
e-
I
h
e-
Ions
Ions
Analytic
Structure Morphology

## Slide 6: SIMS : Secondary ion mass spectrometry
Mass spectrometer
Extraction lens
Sample
Secondary ions
Primary ion beam
Cs+
X-
Imaging
Depth
profiling
Mass
spectrum
~nm
< 100 nm
Isotopic
measurement
Image 3D
HU in D-Mode
Organic S-Mode
Isotopes
6Li, 7Li
High dynamic range
Surface analysis hight sensitivity, hight dynamic range, all elements

## Slide 7: signal intensity
Si Sputter yield:
  He: 0.04
  Ne: 1.3
  O: 1.5
  Cs: 3.3
Sputtering yield typically 20 times lower with He
IionM = T . Yi . db . S . SY . [M]
Cascade of collision
Sputtering rate SY
Ionisation yield  Useful Yield UY
UY
SY

## Slide 8: Ionic Source : State of art
 | Electron impact ion source | Duoplasmatron ion source | Surface ionisation ion source | Hyperion plasma source | Liquid metal ion source | Gas field ionisation ion source | Cold atom ion source
Species | Ar, Xe, O2, SF6, C60, … | Ar, Xe, O2, … | Cs | H, He, O2, Ar, Ne, Kr, Xe | Ga, In, Aun, Bin, … | He, Ne, N2 | Li, Na, Rb, Cs, Ga, …
Virtual source size | 1 mm | 200 µm | 50 µm | 15 µm | 3 nm | ~3 Å | 10 µm
Ion energy for singly charged ions (keV) | 0.1–20 | 0.1–20 | 0.3–20 | 5–30 | 2–30 | 5–45 | 2 – 30
Beam current | 1–1000 µA | 1–1000 mA | 10 µA–100 mA | 1 pA - 1 µA | 1–1000 µA | 1 fA–100 pA | 1 pA – 10 nA
Brightness (A cm-2 sr-1) | 1 | 102 | 102 | 103 | 107 | 109 | >107
Energy dispersion (eV) | 1–10 | 5–20 | 0.2–0.5 | 5–6 | 5–50 | <1 | 0.01 - 10
New FIB platforms
Audinot et al. Reports on Progress in Physics (2021)
SIMS Commercial

## Slide 9: mass spectrometer for FIB platform
Dowsett and Wirtz, Analytical Chemistry (2017)
Usiobo et al. J. of Physical Chemistry (2020)
Quadrupole: no parallel detection capability
×
×
ToF :
  Pulsing required
    Introduction of duty cycle not appropriate for high-resolution high-sensitivity imaging (moreover with low-reactive sources He+, Ne+, ..)
    103-104 times lower sensitivity than DC mode
        analysis time
        sensitivity

## Slide 10: mass spectrometer for FIB platform
Dowsett and Wirtz, Analytical Chemistry (2017)
Usiobo et al. J. of Physical Chemistry (2020)
Quadrupole: no parallel detection capability
×
Magnetic sector: DC mode operation, high transmission, parallel detection
×
Tof :
  Pulsing required
    Introduction of duty cycle not appropriate for high-resolution high-sensitivity imaging (moreover with low-reactive sources He+, Ne+, ..)
    103-104 times lower sensitivity than DC mode
        analysis time
        sensitivity
SOURCE

## Slide 11: focal plane detector for compact SIMS
Magnetic sector with Focal plane detector
Focal plane detector
(FPD)
De Castro et al. Anal. Chem. 2021
Audinot et al. Rep. Prog. Phys. (2021

## Slide 12: Him-sims iNSTRUMENT
TIC : Total Ion Current
Integration of the compact mass spectrometer detector
SIMS detector
SE
He/Ne ion beam column
TIC detector
Wirtz et al. Annu. Rev. Anal. Chem. (2019)

## Slide 13: DUALBEAM-SIMS
Integration of the compact mass spectrometer detector
De Castro et al. Analytical Chemistry (2022)
SIMS
FIB

## Slide 14: Spatial resolution
10 keV Ne, 9.5 pA, FOV 2 um, 1024 px, 300 us
Al+
BAM L200 standard sample
(Certified reference for testing of lateral resolution & length scale calibration)
 Line pair resolution: < 18 nm
Performance
Sijbrandij et al., Microscopy today, 2019
Simulation of 105 impacts
Calculating the distance of sputtered atoms from impact point
Determining FW50
Lateral resolution only dependent on ion-matter
interactions i.e. the collision cascades

## Slide 15: sims system on FIB plateformes
H
H2
6Li
7Li
10B
11B
Na
Si
K
Ti
Imaging
Depth
profiling
Mass
spectrum
SIMS
Ga+ , 3 kV
FOV 3 µm
SE
Si
Ne+ , 25 kV
Ne+, 30 kV

## Slide 16: Photovoltaic solar cells
Localization of Rb Grain, grain boundary (200 ppm)
Photovoltaic solar cells
- CIGS : Rb
Rb concentration map obtained from nano-XRF measurements
P. Schöppe et al., Nano Energy (2017)
CIGS , Cu(In,Ga)Se2, Rb doping
Notes: Solar cells based on Cu(In,Ga)Se2 absorbers are the most efficient ones among all thin film photovoltaics. The current world record efficiency was attained by applying a rubidium fluoride (RbF) post deposition treatment (PDT) to the absorber.
However, it is still not clear why the introduced Rb improves the solar cell performance Localization of Rb ( grain, grain boundary)

## Slide 17: photovoltaic cell
SE
Na
Rb
113In
113In (4.3 %)
115In (96.3 %)
CIGS , Cu(In,Ga)Se2, Rb doping
FOV 20 µm
512 x 512 pixels²
2 ms / pixels
~ 8 min /acq
SE_Rb

## Slide 18: photovoltaic cell
CIGS , Cu(In,Ga)Se2, Rb doping
Detection limit 20 ppm
FOV 5 µm
512 x 512 pixels²
2 ms / pixels
85Rb
113In (4.3 %)
115In (96.3 %)
FOV 20 µm
512 x 512 pixels²
2 ms / pixels
85Rb
Wirtz et al. Annu. Rev. Anal. Chem. (2019)

## Slide 19: FIB section – SIMS analysis
Section view

## Slide 20: photovoltaic cell
Methylammonium lead halides
FAxMAyPbI3
PARTIAL Bromide-Iodide segregation
I. Zimmermann, I et al., J. Mater. Chem. A 2019
J. Usiobo et al., J. Phys. Chem. C 2020
H. Kanda et al., Solar RRL. 2020
P. Gratia et al., J. Am. Chem. Soc., 2016
P. Gratia et al., Acs Energy J. , 2017
(FAPbI3)1-x(MAPbBr3)x

## Slide 21: Imaging with a FPD, HIM-SIMS
Perovskite sample
Cs
Pb
FOV 10 µm
512 x 512 pixels
Ne+ @ 25kV, 3 pA
Dwell time: 250 µs
1 minute
133Cs
206, 207, 208Pb
208Pb
Mass range 7  220
TIC FDP
204Pb
206Pb
207Pb
208Pb
FOV 5 µm
FOV 5 µm

## Slide 22: Life science
Classified as Persistent Organic Pollutants POPs
resistant to environmental degradation typically halogenated organic compounds
Humas
Animas
Air
Water
1. Environmental exposure
2. Bioaccumulation
3. Diversity
Over 3000 PFASs currently on the global market !
Perfluoroalkyl Substances (PFASs)
Courtesy C. Stoffels

## Slide 23: Life sciecnesc
F
CN
FoV : 9 µm
FoV : 30 µm
FoV : 12 µm
Courtesy C. Stoffels
Cs source
ZeroK-SIMS

## Slide 24: correlative microscopy
Combine SIMS with others techniques (EBSD, EDX, BSE, ….)
SIMS + SE / BSE
10 µm
10 µm
Courtesy C. Stoffels

## Slide 25: correlative microscopy
3D + 1 : Photogrammetry & SIMS
Ost et al. Env. Sci. & Tech. 55 (13), (2021)
Image acquisition
3D Morphology / structure
Correction SI Vs Incident angle
Scientific question ( curvature , …)

## Slide 26: Soil science
16O
18O
SIMS
FoV = 5 µm
Seismic decarbonation
exchange in solution
Isotope
Understand how fluid-rock interaction controls the rheology during seismic slip in the presence of fluid phases

## Slide 27: Applications
6Li
5 µm
5 µm
55Mn
De Castro et al.
Analytical Chemistry (2022)
Battery
56Fe
55Mn
7Li
10 µm
10 µm
10 µm
edxCu
edxFe
Nps
SuperAlloy
Notes: Different technologies, detected Li,

## Slide 28: Surface analysis Techniques
TEM
EELS/EDS
AES
LA-ICP-MS
RAMAN
SEM EDS
XPS
XRD
XRF
FTIR
SIMS
Tof-SIMS
µ-XRF
µ-PIXE
100 %
10 %
1 %
0.1 %
100 ppm
10 ppm
1 ppm
100 ppb
10 ppb
1 ppb
100 ppt
0.1 nm
1 nm
10 nm
100 nm
1 µm
10 µm
100 µm
1 mm
1 cm
Nano
Micro
Bulk
ICP-MS
Lateral resolution
Detection limit
Physical limit for 0,3 nm sampling depth
NanoSIMS
TEM
AFM
SEM/HIM
GDMS
TGA/DTA/DSC
ICP Techniques
GC-MS
Optical
Spatial resolution
Incident Particle
X Ray
X Ray
h
h
e-
I
h
e-
Ions
Ions
Analytic
Structure Morphology
FIB-SIMS

## Slide 29: Conclusions
SIMS capability on Focused Ion Beam - based platforms
SE imaging combined with a sub-20 nm SIMS imaging
Application in material, soil, and life science
Continue detector for double focusing MS
Cryo stage @ 130 K, Cryo Transfer UHV
Correlative microscopy ( EDX, STEM, ….)
jean-nicolas.audinot@list.lu